## Slide 1: 7.3.4 Using power amplifiers as high current buffers in low Z VNA measurements
Presented by:
Jithendra H S

## Slide 2: What?
Low Z(PDN impedance):
  Impedance in terms mΩ
  Package leads
  Ceramic capacitors
  On-chip capacitance
  Power and ground planes
  VRM impedance
Impedance measuring methods:
  Conventional(DMM)
  4 wire(DMM)
  1-port(VNA)
  2-port(VNA)
Conventional
4-Wire
1-port
2-port

## Slide 3: Why?
2-Port Method:
The current from port 1 is limited by the 50-ohm resistance of the VNA source.
This current flows through the low impedance of the DUT, generating a very small voltage across the DUT.
The voltage across the DUT is launched into port 2, where it is received and measured.
The voltage ratio, S21​ (voltage at port 2 to the source voltage from port 1), provides information about the DUT's small impedance.
Current-Amplified 2-Port method:
When the measured voltage level is near the noise floor of the VNA, it becomes challenging to accurately differentiate the low impedance of the DUT.
To overcome this, an amplifier is introduced to:
  Boost the voltage, which increases the current across the DUT.
  Create a larger voltage drop across the DUT, making it easier for port 2 to measure the impedance.

## Slide 4: How?
Know the measuring device features and limitations:
  Bandwidth:- Measuring amplitude at selected frequency
  Noise floor:- Measuring S21 by keeping port 2 open
  Power level:- Maximum operating power of ports
Settings for measurements:
  1KHz-1GHz Bandwidth
  Logarithmic sweep(1000 points)
Test known low Z components for correctness:
  Characterization: Measurement
  Analysis: Simulation
Selection of power amplifier
  Constant insertion loss within the Bandwidth
  Low output impedance
Which is better?
  Pre-amp
  Post amp

## Slide 5: Timeline
Week | Tasks | Remarks
1 | Introduction | Completed
2 & 3 | Understanding the VNA | Completed
4 & 5 | Selecting the Power amplifier, look for better options | Completed
6 & 7 | Results interpretation and Model simulation | On-track
8 & 9 | Report preparation and corrections | 
10 | Buffer!!! | 
Week | Tasks | Remarks
6 & 7 | Introduction | 
8 & 9 | Characterize set of existing board, Update board design(spin 1) | 
10 & 11 | Characterize spin 1 boards, Update board design(spin 2) | 
12 & 13 | Results interpretation and Model simulation | 
14 & 15 | Report preparation and corrections | 
16 | Buffer!!! | 
Project 1:
Low Z measurement
Project 2:
USB Signal sniff boards

## Slide 6: Week 3: Understanding the VNA
Port 1 output:
  Max power output : -10dBm
  Low power output: -17dBm
  Not ideal sine wave : due to frequency sweep
200mVpp at High Z
100mVpp at 50ohm
50mVpp at 50ohm with a -6dB attn
The voltage reduced to half when the oscilloscope channel was switched from high impedance to 50Ω. According to the voltage divider circuit, half the voltage drop occurs across the load resistance only when it is equal to the internal resistance. Therefore, the internal resistance is equal to 50Ω.

## Slide 7: With High BW scope
1KHz: attenuated
10KHz: ~100mVpp
50MHz: ~100mVpp
1GHz: Not sure of what is happening

## Slide 8: Noise floor
Measured port 2 without any connection after calibration
Noise floor is not constant over the frequency, it varies with frequency as:
  -90dB up to 100MHz
  -70dB up to 500MHz
  -50dB above 500MHz

## Slide 9: Noise floor- Log sweep
~-100dB
So, with -100dB noise floor the minimum impedance that VNA can measure is 250uΩ

## Slide 10: Power level
Port 2 safe input level:
Max input power : 0dBm(225mV)

## Slide 11: Measured known 10mohm impedance

## Slide 12: Amp responses
Attenuated by -30dB attenuator
Attenuated by -20dB attenuator

## Slide 13: Week 4: RF amplifier gain
Gain= 27.9±0.2dB
%
=

## Slide 14: Week 4: RF amplifier gain
Gain= 27.9±0.2dB
%
=

## Slide 15: Input and Output Z
Zout =
Zin = ~90ohm
Zout = ~35ohm

## Slide 16: Noise floor with RFamp
20uohm
We can measure the low Z impedance up to:
Therefore 20uΩ
Floating
Safety margin after considering noise floor:
SDR VNA: ~95dB, RF amp : ~ 93dB
-93dB –(+28dB) = -121dB -> 22.2u Ω
121dB

## Slide 17: 10m Ohm resistor
Expected to be 10mΩ, but measured:
VNA: 11mΩ
Turbo VNA: 12mΩ

## Slide 18: Via resistance – layer 2
Rule of thumb:
Via height = length = 10mil
Via thickness = 0.5Oz
Via diameter = 13mil
R per Sq = 1mΩ
N of Sq = 10/(3.14 *13) = 0.244
R = 0.244* 1mΩ = 244uΩ
The 289uΩ measured is in close agreement with the rule of thumb

## Slide 19: Via resistance – layer 2
Rule of thumb:
Via height = length = 10mil
Via thickness = 0.5Oz
Via diameter = 26mil
R per Sq = 1mΩ
N of Sq = 10/(3.14 * 26) = 0.122
R = 0.122* 1mΩ = 122uΩ
The 129uΩ measured is in close agreement with the rule of thumb
Turbo VNA stands out, as we can see we can measure well below the noise floor of VNA(250uΩ)

## Slide 20: Via resistance simulation– layer 2
Rough estimate:
From ADS Via model:
R= 122 to 130uΩ

## Slide 21: Via resistance simulation– layer 2
Putting the numbers:
From ADS Via model:
R= 77 to 80uΩ Why is this different than rule of thumb and actual measurement?
- Will be discussed next time

## Slide 22: So What?
Measuring the ultra low Z with low cost VNA and low-cost RF amp
Methodology for characterizing the RF amp for low Z measurement

## Slide 23
Questions?

## Slide 24: Dummy

## Slide 25: 7.3.4 Using power amplifiers as high current buffers in low Z VNA measurements(Turbo VNA)
Presented by:
Jithendra H S
3/26/2025

## Slide 26: Agenda
Problem
Proposed Solution
About RF amplifier
Measurement analysis
Limitation

## Slide 27: Problem
*Simulated using ADS
-100dB VNA Noise floor
Via response
Min 250uΩ
Via resistance 75uΩ

## Slide 28: Solution
-100dB VNA Noise floor
Via response
-130dB TURBO VNA Noise floor
Min 250uΩ
Via resistance 75uΩ
Extended resistance up to 8uΩ
Normal VNA setup
Turbo VNA setup

## Slide 29: RF amp Improvements
3.3uF DC blocking cap(latest)
0.1uF DC blocking cap(early)
Gain = +30.927dB

## Slide 30: Revised Noise floor
-450uΩ
-14uΩ
-95dB
-125dB
Moving avg: 10pts

## Slide 31: Why Post amp?
Preamplification
Post amplification

## Slide 32: Test Board
Up to layer 2
Up to layer 2
13mil dia
26mil dia
13mil dia
26mil dia
Port extension
10m ohm ±1%

## Slide 33: 10m Ohm resistor
10mΩ±1%
11mΩ
8mΩ

## Slide 34: Phase of 10m Ohm resistor
Z = R + jwL
Failing to sample the phase properly

## Slide 35: Attenuator setup
-30dB
-30dB
-20dB
-10dB
-6dB
-3dB
-1dB

## Slide 36: Consistency check
5°
50°
70°
50°
Supporting resource from Tim Wang Lee: https://groups.io/g/VNWA/topic/looking_for_ways_to_improve/74472949

## Slide 37: Via - layer 2
200uΩ
226uΩ
Rule of thumb:
Via height = length = 10mil
Via thickness = 0.5Oz
Via diameter = 13mil
R per Sq = 1mΩ
N of Sq = 10/(3.14 *13) = 0.244
R = 0.244* 1mΩ = 244uΩ
Conclusion: Measured values are around 11% off

## Slide 38: Via- layer 2
75uΩ
83uΩ
Rule of thumb:
Via height = length = 10mil
Via thickness = 0.5Oz
Via diameter = 26mil
R per Sq = 1mΩ
N of Sq = 10/(3.14 * 26) = 0.122
R = 0.122* 1mΩ = 122uΩ
*Moving avg: 10pts
Conclusion: Measured values are around 10% off

## Slide 39: Summary table
DUT | SDR VNA |  | SDR Turbo VNA | 
 | Measured | Error % | Measured | Error %
10mΩ | 11mΩ | 10% | 8mΩ | 20%
Via-GL2-13mil(200uΩ) | Saturated | - | 226uΩ | 13%
Via-GL2-26mil(75uΩ) | Saturated | - | 83uΩ | 10.6%

## Slide 40: So What?
Measuring ultra-low impedance using an affordable VNA and RF amplifier.
Approach for characterizing the RF amplifier for low impedance measurements.
Despite phase shift limitations at lower frequencies, we can still approximate the structure's resistance closely.

## Slide 41
Questions?

## Slide 42: 7.3.4 Using power amplifiers as high current buffers in low Z VNA measurements(Turbo VNA)
Presented by:
Jithendra H S
4/10/2025

## Slide 43: Agenda
Problem statement & Proposed Solution
Artifacts and Improvements
Measurement analysis
Conclusion

## Slide 44: Problem Statement
*Simulated using ADS
-100dB VNA Noise floor
Via response
Min 250uΩ
Via resistance 75uΩ

## Slide 45: Solution
-100dB VNA Noise floor
Via response
-130dB TURBO VNA Noise floor
Min 250uΩ
Via resistance 75uΩ
Extended resistance up to 8uΩ
Normal VNA setup
Turbo VNA setup

## Slide 46: RF amp Improvements
3.3uF DC blocking cap(latest)
0.1uF DC blocking cap(early)
Gain = +30.927dB

## Slide 47: Common current artifact
*Source: HSDD wk 3 lecture slides

## Slide 48: Revised Noise floor
Moving avg: 10pts

## Slide 49: Why Post amp?
Preamplification
Post amplification

## Slide 50: Test Board
Up to layer 2
Up to layer 2
13mil dia
26mil dia
13mil dia
26mil dia
Port extension
10m ohm ±1%

## Slide 51: 10mΩ resistor
11mΩ
10.65mΩ
8mΩ

## Slide 52: Similarity metric
Nominal value of the resistor is 10mΩ±1%
The simulation showed the resistance value is about 10.65±.5mΩ
Turbo VNA measurement is 8mΩ, which is off by 20-25%

## Slide 53: Via - layer 4
1mΩ
1mΩ
1mΩ
Thickness of Via: 1.2mil

## Slide 54: Via - layer 2
Rule of thumb:
Via height = length = 8mil
Via thickness = 1.2mil
Via diameter = 13mil
R per Sq = 0.5mΩ at 1Oz
N of Sq = 8/(3.14 *13) = 0.195
R = 0.195 * (1.4/1.2)0.5mΩ = 113uΩ
179uΩ
192uΩ
122uΩ
*Working on finding the Uncertainty

## Slide 55: Via- layer 2
Rule of thumb:
Via height = length = 8mil
Via thickness = 1.2mil
Via diameter = 26mil
R per Sq = 0.5mΩ at 1Oz
N of Sq = 8/(3.14 * 26) = 0.097
R = 0.097 *(1.4/1.2) 0.5mΩ = 56uΩ
*Moving avg: 10pts
58uΩ
100uΩ
68uΩ
*Working on finding the Uncertainty

## Slide 56: So What?
10mΩ
Via layer 4
Via layer 2
VNA NF
Turbo VNA NF
Measured ultra-low impedance using an affordable VNA and RF amplifier.
  10m ohm, Via etc
Came up with an approach to characterize the RF amplifier for low impedance measurements.
  I/P, O/P Z, Wide band gain,
Reduced the limitations, resulting in a wide band amplifier and very less common current
Despite the gain at lower frequencies, we can still approximate the structure's resistance better than a first-order approximation
Common current artifact

## Slide 57
Questions?

## Slide 58: Dummy

## Slide 59: Back up

## Slide 60: Simulated Pre and Post amp

## Slide 61: Input and Output Z of RF amp

## Slide 62: Week 6 update
  Analysing the new set of measurements in terms of scaling the RF amp output be removing the influence of 30dB attenuators used to measure Gain using ADS(5hrs)
  Compared Pre and Post amplification methods on 10mOhm resistor(4hrs)
  Characterized all other low cost RF amps exist in the lab(2hrs)
Week 7:
Take new measurements with respect to Via(2layer and 4 layer) (2hrs)
Characterizing the Via impedance wrt to pre and post amplification method(8hrs)

## Slide 63: Week 7 update
  Took new measurements with respect to Via(2layer and 4 layer) (1hrs)
  Characterizing the Via impedance wrt to pre and post amplification method(8hrs)
Week 8:
Make sense of the uncertainties in pre and post at lower frequencies(6-8 hrs)
Starting to measure the and verify the passive sniffing boards(2hrs)

## Slide 64: Week 8 update
Compared VNA and Turbo VNA measurements using ADS(4hrs)
Tried to analyse the mismatch at the lower frequencies (6hrs)
Week 9:
Draft the article and come up with a story(5 hrs)
Recollect the data if required for the article for improvisation(4hrs)

## Slide 65: Week 10 update
Re-did the measurements in the post-amp method and bring them into ADS(4hrs)
Tried to analyze the mismatch at the lower frequencies (6hrs)
Week 11:
Draft the article and come up with a story(7 hrs)

## Slide 66: Week 12 update
Research about the cause of phase change artifact(6hrs)
Re-did the measurements and analyzed them in ADS(4hrs)
Week 13:
Rework the draft paper and add new findings(6hrs)

## Slide 67: Week 13 update
Created a new low impedance structure with the help of components available in the lab, analyzing the results(6hrs)
Rework the draft paper and added new findings(2hrs)
Week 14:
Need to find the DC resistance of the Via and the resistance
Analyse the new structure measurements with DC resistance