## Slide 1: 7.3.4 Using power amplifiers as high current buffers in low Z VNA measurements
Presented by:
Jithendra H S
05/01/2025

## Slide 2: Introduction
The objective of this project is to overcome limitations and measure ultra-low impedance structures, such as resistors, vias, and PDN impedance, using a low-cost SDR VNA and an RF amplifier.
Measurement setup:

## Slide 3: Road map
VNA Characterization
Why an amplifier?
Amplifier optimization
Common current artifact
Ultra low Z measurement

## Slide 4: VNA Overview
Power: Max = -10dBm
Min = -17dBm
Voltage Output:
Vpp=200mV
Noise floor:
-100dB up to 100MHz
-70dB up to 500MHz
-50dB above 500 MHz
Max Input power:
0dBm -> 225mV
With a -100dB(1uV) noise floor, the minimum impedance that SDR VNA can measure is up to 250uΩ

## Slide 5: VNA config & Test vehicle
Up to layer 2
Up to layer 4
13mil dia
26mil dia
13mil dia
26mil dia
Port extension
calibration
10m ohm ±1%
Known structure
Parameter | Settings
Sweep range | 1 KHz to 200 MHz
Datapoints | 1000
Sweep time | 30s

## Slide 6: Common current artifact
Due to the common current
Ferrite choke
Common current limits the low Z measurement to ~25mΩ
DUT

## Slide 7: Limitation of Normal VNA
Rule of thumb:
Via height = length = 8mil
Via thickness = 1.2mil
Via diameter = 26mil
R per Sq = 0.5mΩ at 1Oz
N of Sq = 8/(3.14 * 26) = 0.097
R = 0.097 *(1.4/1.2) 0.5mΩ = 56uΩ
-100dB VNA Noise floor
Via response
Min 250uΩ
Via resistance 56uΩ
*Simulated using ADS
-110dB(316nV) measurement is not at all possible to measure with a Normal VNA, if you measure always result in 250uΩ
26 mil Dia

## Slide 8: Why Amplification?
+20dB
-100dB VNA Noise floor
Via response
Via response
-100dB VNA Noise floor
Adding gain to the voltage drop across DUT helps the VNA port to measure the signal without being limited by the noise floor
After that, we need to vector rescale the amplified insertion loss using ADS to get the actual insertion loss across the DUT
Turbo VNA setup
Normal VNA setup

## Slide 9: Revised noise floor
*Moving avg: 10pts
Due to the addition of both ferrite chokes & the RF amp, the effective noise floor results in measuring impedance up to 8uΩ
Kept Open
Insertion loss
Corresponding Impedance
Turbo VNA setup

## Slide 10: 10mΩ resistor
10.65mΩ
7.65mΩ
Measuring known structures will reveal the uncertainty in measurement between the Normal and Turbo VNA.
Insertion loss
Corresponding Impedance
RL model
DUT
---Turbo_VNA_meas
---Normal_VNA_meas
---Simulated

## Slide 11: Similarity metric
Nominal | DC 4 wire method | Normal VNA | Uncertainty | Turbo VNA | Uncertainty
10mΩ±1% | 10.618mΩ | 10.65mΩ | ±0.04mΩ(0.03%) | 7.675mΩ | ±2.9mΩ(27%)
Normal_VNA_meas - simulated
Turbo_VNA_meas - simulated

## Slide 12: Via - layer 2 -13mil dia
Rule of thumb | ADS Hacked | DC 4 wire method | Normal VNA | Uncertainty | Turbo VNA | Uncertainty
113uΩ | 230uΩ | 568uΩ | 179uΩ | ±51uΩ(22%) | 192uΩ | ±38uΩ(16%)
*Moving avg: 10pts
±300uΩ
~550uΩ
230uΩ
179uΩ
192uΩ
Impedance
Similarity metric
RL model
DUT

## Slide 13: Via- layer 2 – 25mil dia
*Moving avg: 10pts
Rule of thumb | ADS Hacked | DC 4 wire method | Normal VNA | Uncertainty | Turbo VNA | Uncertainty
56uΩ | 90uΩ | 346uΩ | NA | NA | 68uΩ | ±22uΩ(24%)
90uΩ
~350uΩ
68uΩ
Overlap with Noise floor
???
±300uΩ
Impedance
Similarity metric
RL model
DUT

## Slide 14: Summary from multiple responses
10mΩ
Via layer 4
Via layer 2
VNA NF
Turbo VNA NF
Common current artifact
Still see a common current artifact with the test vehicle at 0.5mohm range.
Able to extend the noise floor of VNA from -100dB up to -130dB, as we can see with Turbo VNA NF.
By scaling the difference in gain from Normal and turbo VNA, we can reduce the uncertainty and improve the Turbo VNA response.

## Slide 15: Takeaway
Can measure ultra-low impedance of up to ~8uΩ using an affordable VNA and RF amplifier.
Steps to characterize the RF amplifier for low impedance measurements in terms of I/P Z, O/P Z, and Wideband gain.
Reduced the limitations:
Common current: Ferrite chokes enabled low flat resistance measurement.
  14mΩ - > 0.45mΩ
High pass filter: Change in DC blocking caps enabled low-frequency sweeping.
  100KHz - > 1KHz
Despite the gain at lower frequencies, we can still approximate the structure's resistance of ~90uΩ with an uncertainty of ~24%, which is better than a first-order approximation & Normal VNA.

## Slide 16: Thank You

## Slide 17
Questions?

## Slide 18: Why Post amp?
Preamplification
Post amplification
-30dB attenuator as DUT
-30dB attenuator as DUT

## Slide 19: Via - layer 4
1mΩ
1mΩ
1mΩ
Thickness of Via: 1.2mil

## Slide 20: Test Vehicle artifact
Manually measured 280uohm

## Slide 21: Via - layer 2
179uΩ
192uΩ
122uΩ
Rule of thumb | ADS | DC 4 wire method | Normal VNA | Uncertainty | Turbo VNA | Uncertainty
113uΩ | 122uΩ | 568uΩ | 179uΩ | ±57uΩ(46%) | 192uΩ | ±70uΩ(57%)
~550uΩ
*Moving avg: 10pts

## Slide 22: Via- layer 2
*Moving avg: 10pts
58uΩ
???
68uΩ
~350uΩ
Rule of thumb | ADS | DC 4 wire method | Normal VNA | Uncertainty | Turbo VNA | Uncertainty
56uΩ | 58uΩ | 346uΩ | NA | NA | 68uΩ | ±10uΩ(17%)
Overlap with Noise floor